## GTEM
Voltage ratio | 0.00099651 | Vout/Vin | Based on your input voltage divider ratios | [5 empty] | L Line Input Impedance | 1000 | Ohm
Current ratio | 0.014563 | Vout/Ain | Based on your input current shunt/CT ratios | [5 empty] | V Channel Input Impedance | 400 | kOhm
Impulse converter
MMode | [13 empty] | 973 | W
Lgain= | 100: Gain 1 |  |  |  | 100 |  |  |  |  | 1000000000 |  |  |  | 3.6999 | seconds per impulse | 0.27028 | Hz
Ngain= | 10: Gain 1 |  |  |  | 10 | [8 empty] | 973 | impulses in one hour | 0.21978 | Hz
LNSel= | 1: Meter on L line |  |  |  | 1
DisHPF= | 00: HPF1 and HPF0 |  |  |  | 00
Amod= | 0: Fwd or rev energy pulse output |  |  |  | 0
Rmod= | 0: Fwd or rev reactive energy pulse output |  |  |  | 0
Zxcon= | 10: All zero crossing |  |  |  | 10
Pthresh= | 0010: 3.125% |  |  |  | 0010
Mmode Reg= | 1001010000100010 | [10 empty] | 0.11257
= | 10010100 | 00100010
= | 94 | 22
Mmode Reg= | 0x9422
PL Constant
UN= | 300 | V | Reference voltage, unit is V
Ib= | 12 | A | Basic current, unit is A
GL= | 1 | V/V | L line current circuit gain
VL= | 298.9537 | mV | Sampling voltage of the L line circuit at Ib, unit is mV
VU= | 174.7573 | mV | Sampling voltage of the voltage circuit at UN, unit is mV
MC= | 1000 | imp/kWh | Pulse constant of the energy meter, unit is imp/kWh
PL_Constant= | 12173811
= | 00B9C1F3
PLconstH(0x21)= | 0x00B9
PLconstL(0x22)= | 0xC1F3
Small-Power Mode and Power Offset Compensation
1. Disconnect the current circuit of the energy meter;
2. Write '0xA987' to the SmallPMod register (04H) to enter small-power mode
3. Read out L / N line active / reactive power in small-power mode many times to get a mean value;
4. Write the complement of the above mean value to the corresponding L / N line active / reactive power offset registers;
5. Write any non-0xA987 value to the SmallPMod register (04H) to exit small-power mode
L Line Gain Calibration (Calibrate PL Constant Energy Metering)
L line gain calibration is performed when power factor PF=1.0 and the current is Ib
ETSA Frequency | 0.269 | Hz | Pulse output frequency as measured from ETSA metering using photodiode.
My Frequency | 0.219 | Hz | Pulse output frequency as measured from my energymeter using photodiode.
Error % = | -18.5874 | %
Lratio = | 0.22831 |  |  | 65536 |  | 1010 | 10 | 5
Lgain= | 7481 |  | 58054 |  |  | 0101
Lgain= | 0x1D39
Example Values
Voltage RMS Gain and L Line Current RMS Gain | [7 empty] | Temp |  |  |  |  | 8V AC RMS In
Vol_mea= | 159.1 | V |  | Readout value of the Urms register (49H) |  |  |  | 159.1 | 0.0060265 |  |  |  | 158.4
Un= | 244.1 | V |  | Actual Voltage |  |  |  | 244.1 | 0.24325 | 40.3631 |  |  | 246
Ugainold= | 26400 |  |  | 26400 is the default value | [8 empty] | _lgain = 0x1D39; // Use XLS to calculate these values, PL CONSTANT. Examples: 0x1D39;
_ugain = 0xA028; // Use XLS to calculate these VOLTAGE RMS values. Examples: 8V 0xA028 ¦ 12V 0x9F9A or 0x9E38
_igain = 0x7160; // Use XLS to calculate these CURRENT GAIN values. Examples: 0x7160;
_crc1 = 0xAE70; // Important! Run this application, then take auto CRC1 calculated values and update here. Examples: 0xAE70
= | 40504 | [11 empty] | _crc2 = 0xDC3E; // Important! Run this application, then take auto CRC2 calculated values and update here. Examples: 8V 0xDC3E ¦ 12V 0x51AF or 0xF24C
Ugain= | 0x9E38 |  |  | Will require CRC2 to be updated (GTEM-1_Defaults)
Cur_meaL= | 1.6 | A |  | Readout value of the Urms register (49H) |  |  |  | 1.6
Ib= | 1.486 | A |  | Actual Current |  |  |  | 1.486 |  |  |  |  | 12V AC RMS In
Igainold= | 31251 |  |  | 31251 is the default value | [8 empty] | 159.1
244.1
_lgain = 0x1D39; // Use XLS to calculate these values, PL CONSTANT. Examples: 0x1D39;
_ugain = 0x9F37; // Use XLS to calculate these VOLTAGE RMS values. Examples: 8V 0xA028 ¦ 12V 0x9F9A or 0x9E38
= | 29024 |  |  | Based on a YHDC SCT-013 100A/50mA Current Transformer | [8 empty] | _igain = 0x7160; // Use XLS to calculate these CURRENT GAIN values. Examples: 0x7160;
Igain= | 0x7160 |  |  | Will require CRC2 to be updated (GTEM-1_Defaults) |  |  |  |  | offset=0.003A |  | IgainL | 340 | _crc1 = 0xAE70; // Important! Run this application, then take auto CRC1 calculated values and update here. Examples: 0xAE70
[12 empty] | 154 | _crc2 = 0xFC4C; // Important! Run this application, then take auto CRC2 calculated values and update here. Examples: 8V 0xDC3E ¦ 12V 0x51AF or 0xF24C
The current offset calibration should be performed at reference voltage and no current.
Calibration is performed by reading the current, multiplying it with the above Igain/2^16 and 2^8, | [12 empty] | The above example values are what I have used on GTEM to test. Based on YHDC SCT-013 100A/50mA Current Transformer and DAT01A Mains Transformer
calculating the complement and writing the result to the L line current offset register (IoffsetL, 35H). | [12 empty] | These WILL vary from one setup to another due to tolerances or the components, clamp and transformer.
Dave Williams, DitroniX
Calibration of voltage offset is similar to that of current offset, but voltage offset
calibration is not applicable for self-powered meter
Voltage Sag Setpoint
Un= | 230 | Nominal voltage
Ugain= | 40504 | From Voltage RMS Gain calc
x%= | 0.94 | 6% below nominal
SagTh = | 6183
SagTh = | 0x1827
PL_Constant=int(838860800×(G_L×V_L×V_U)/(MC×U_(N`)×I_b ))
U_gain=int((Ugain_old×U_n)/(Vol_mea ))
I_gain=int((Igain_old×I_b)/Cur_mealL )
SagTh=(100×U_n× √2×𝑥%)/(4×𝑈_𝑔𝑎𝑖𝑛/32768)
Note on A5: Wendland, Ryan: L line current gain
Note on A6: Wendland, Ryan: N line current gain
Note on A7: Wendland, Ryan: This bit specifies metering as L line or N line when metering mode is set to flexible mode by MMD1 and MMD0 pins.
Note on A8: Wendland, Ryan: These bits configure the High Pass Filters (HPF) after ADC.
Note on A9: Wendland, Ryan: CF1 output for active power
Note on A10: Wendland, Ryan: CF2 output for reactive power
Note on A11: Wendland, Ryan: These bits configure zero-crossing mode. The ZX pin outputs 5ms-width high level when voltage crosses zero
Note on A12: Wendland, Ryan: These bits configure the L line and N line power difference threshold in anti-tampering mode